## Slide 1
Team should train/practice &
have necessary supplies at bedside
Clamp drainage & return lines
Call for help
Return to pre-ECMO vent settings
Exchange of oxygenator or entire circuit may be needed
by Nick Mark MD & Jonah Rubin MD
ecmo troubleshooting
onepagericu.com
@nickmmark
@JonahRubinMD
Link to the most current version →
ONE
PUMP
OXYGENATOR or
MEMBRANE LUNG
CC BY-SA 3.0 v1.0 (2021-03-14)
37
100
60
21
40
FiO2 (%)
Flow (lpm)
60
55
45
40
35
30
100
60
50
CIRCUIT PROBLEMS
PUMP PROBLEMS
MEMBRANE/OXYGENATOR PROBLEMS
GAS EXCHANGE PROBLEMS
CHATTER is visible shaking of the tubing from variable venous drainage when the pump is trying to drain more than what native venous return allows. Pump may repeatedly start/stop causing CUT OUTS. Options: Reduce RPM, administer fluids, assess cannula positioning
CIRCUIT THROMBOSIS: Clots in the tubing can increase flow resistance. Clots in oxygenator may also impair gas exchange. Acutely or over time, this may lead to OXYGENATOR FAILURE. Anticoagulation may prevent clot formation.
CANNULA DISPLACEMENT compromises the circuit and can cause an air embolism, major bleed, and reduced blood flows. Follow ECMO emergency protocol and consider massive transfusion.
AIR EMBOLISM can cause immediate pump, oxygenator, or circuit failure. Follow ECMO emergency protocol and place head down. Locate source to prevent further entrainment. If venous source, may place in left lateral position and aspirate from drainage cannula
PUMP FAILURE may occur from CIRCUIT THROMBOSIS, AIR EMBOLISM, or other insult. Follow emergency protocol . A backup pump should be readily available. With some ECMO products, a mechanical crank can be used temporarily until a new pump is available.
INADEQUATE FLOW can occur from kinking, cannula malposition, decreased venous return (± CHATTER), and CIRCUIT THROMBOSIS (same RPM yielding less LPM flow).
Approach: Treat the cause! Some patients, especially with high BMI, may need an extra venous drainage cannula to achieve goal flow.
RECIRCULATION occurs when drainage cannula draws oxygenated blood from the return cannula, reducing the efficiency of the circuit. Increased SpO2 and red coloration of venous drainage may be noted.
Approach: Reposition/reconfigure cannulas, reduce flow (if tolerable)
HYPOXEMIA is common as oxygenated blood always mixes with deoxygenated
native circulation, although this may not reflect insufficient oxygen delivery. A low SpO2 may be tolerable, however SpO2 < 85% may be critical.
Options:
1. Increase flow – may be limited by CHATTER, diminishing returns if flow >7LPM
2. Increase FDO2 – diminishing returns when post-oxygenator SpO2 is ~100%
3. Reduce native cardiac output (e.g. beta-blocker) – does not change oxygen content but may increase SpO2
4. Last resorts: Consider transfusions to increase CaO2 (given shunt physiology),
adding a second circuit in parallel, or recruiting native lung with ventilator
HYPERCAPNIA/ACIDOSIS can worsen with increased CO2 production (e.g. fever) or reduced clearance (worsening function of native lung or membrane lung), or impaired renal compensation.
Approach: Increase SWEEP (similar to minute ventilation) for acute correction, & correct the underlying cause (fever control, RRT, etc.)
OXYGENATOR FAILURE may occur from AIR
EMBOLISM, CIRCUIT THROMBOSIS (most common), or any serious mechanical insult. Follow the ECMO emergency protocol to exchange oxygenator.
Clamps are used emergently to stop blood flow
50cc syringes can be used to remove air from the circuit
Vent ports are used to purge the oxygenator of air in an emergency. Tilt & open vent to purge.
RISING TRANSMEMBRANE PRESSURE: At constant flow, this indicates worsening resistance (usually thrombosis) & therefore worsening function, of the oxygenator. Check a post-membrane ABG to assess oxygenation. Ultimately may require oxygenator exchange.
POST-CANNULATION SYNDROME is a SIRS-like inflammation as blood reacts to the oxygenator at 24-48 hrs of initiation. Options: Monitor, consider steroids, in cases with vasoplegia requiring high dose pressors consider methylene blue
Recirculation (%)
“Normal” recirculation is up to ~25%)
ECMO EMERGENCY PROTOCOL
Notes: Stepwise approach for Chugging/Chatter
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC6026563/
What is the circuit
Indications (EOLIA + others?)
Has also been used lower flow, higher sweep (ECCO2R)
Series vs parallel physiology
Shunt physiology