## Slide 1
by Nick Mark MD
intra-aortic balloon pump
onepagericu.com
@nickmmark
Link to the most current version →
ONE
v1.0 (2022-05-05) CC BY-SA 3.0
HELIUM
Pressure transducer & pressure bag used to measure pressure within the aorta. Some newer IABP systems use a fiberoptic sensor to measure pressure changes more rapidly.
Helium tank
Helium is used because it has lower airflow resistance, which enables faster balloon inflation & deflation. Helium is also highly soluble in blood, so it is less likely to cause air embolism if the balloon ruptures.
Sheath side port can be used for blood draws
Radio-opaque tip should be ~2cm distal to the left subclavian
Balloon should be proximal to the renal arteries
Sheath is placed in artery & allows balloon catheter to be advanced to the aorta
Sterile sleeve permits adjusting the depth of the intra-aortic balloon
Unassisted diastolic
Unassisted systolic
UNASSISTED
UNASSISTED
ASSISTED
EKG
Aortic Pressure
Balloon Pressure
Balloon inflation
Time
Augmented diastolic pressure is increased by balloon inflation
Assisted diastolic pressure decreased by balloon deflation
The controller unit rapidly inflates the intra-aortic balloon just after aortic valve closure
A 25-50 ml aortic balloon fills 85-90% of the thoracic aorta
Intra-aortic balloon counter-pulsation is a method of invasive hemodynamic support. A catheter is placed through an arterial sheath and advanced into the thoracic aorta. A gas filled balloon at the end of the catheter inflates in sync with the cardiac cycle:
· Balloon inflates in diastole  increases coronary perfusion
· Balloon deflates in systole  decreases afterload & increases LV stroke volume (SV)
The thick-walled LV is only perfused by the coronaries during diastole. By increasing diastolic pressure, the IABP improves coronary artery perfusion.
Intra-aortic balloon pumps (IABP) can be used to support people in cardiogenic shock, those undergoing revascularization, or as a bridge to to intervention or for interfacility transport. It can also be used as an LV vent in patients receiving VA ECMO.
While ithas salubrious hemodynamic effects, neither RCTs nor meta-analysis has found a survival benefit to IABP use in people with cardiogenic shock.
Assisted systolic pressure is decreased
PRINCIPLE:
AORTIC
VALVE
OPENS
LV volume
LV pressure
UNASSISTED
IABP ASSISTED
AORTIC VALVE CLOSES
LVSP = Systolic BP
LVEDP
Proper IABP support depends on precise inflation & deflation timing. In order to sync the balloon with the cardiac cycles, the controller uses EKG or aortic pressure to trigger inflation. Asynchronous (e.g. set at a rate of 80) can be used as a backup.
· EKG: triggers inflation at the middle of the T-wave & triggers deflation at the peak of the R wave. Arrythmias such as afib & pacer spikes can disrupt EKG triggering. (Atrial pacing is particularly problematic). ECG triggering is preferable in most patients.
· Pressure: aortic pressure triggers inflation at the dicrotic notch and deflates based elapsed time. Pressure trigger is inherently less precise than EKG (the pressure wave propagates slower than electricity)
ECG electrodes are used to trigger inflation
TRIGGER:
A sharp V morphology at the dicrotic notch indicates optimal timing
7.5Fr
catheter
Stroke work (area of the PV curve) is similar
Stroke volume is increased with IABP support
SV
Ideally the balloon would inflate immediately at the onset of diastole (40 msec before the dicrotic notch). Modern IABP have automatic timing but timing can (& should be) manually optimized.
Consequences of improper timing include:
· Early balloon inflation increased afterload
· Late balloon inflation  decreased diastolic augmentation
· Early balloon deflation  no decrease in myocardial O2 demand
· Late balloon deflation  increases afterload
· Poor diastolic augmentation  suboptimal coronary perfusion
Timing should always be adjusted in 1:2 support mode.
TIMING:
ANTICOAGULATION:
The catheter is potentially thrombogenic, however based on limited data, routine anticoagulation may not be required in 1:1 mode. It may be necessary if augmentation is reduced to 1:2, 1:3.
· Limb ischemia  monitor distal pulses
· Juxta-renal positioning  monitor urine output, CXR
· Bleeding, hematoma  monitor sheath site, check coags
· Decreased augmentation  consider low SVR (sepsis), decreased cardiac output (AIBP requires a CI of >1 to “augment”)
· Worsening cardiac ischemia  adjust timing/trigger
COMPLICATIONS & MONITORING:
AUGMENTATION/WEANING:
The level of hemodynamic support can be adjusted. By default, the IABP augments every cycle (1:1). It can be decreased to every other (1:2) or every third (1:3) cycle. Timing & adjustments should be done in 1:2 mode. Decreasing support (1:1  1:2  1:3) is often done to wean IABP, though the actual coronary perfusion provided in 1:3 may be minimal.
IABP 1:2 AUGMENTATION