## Slide 1
weak signal
· ↓↓ perfusion index
· try alternate positions/probes
by Nick Mark MD
pulse oximetry
onepagericu.com
@nickmmark
Link to the most current version →
ONE
PULSE OX LAG TIME
The pulse oximeter probe can be applied to different locations.
· Finger – most accurate; thumb may be more accurate than other fingers. Uses a transillumination method to measure.
· Ears & Forehead – slightly less accurate but may be more reliable in vasoconstricted states or hypothermia; uses backscatter method to measure.
No probe type is clearly superior; trial & error may identify the best probe type for an individual patient.
Red LED
IR LED
100%
80%
60%
40%
20%
0%
DeoxyHb
OxyHb
Absorption (%)
Wavelength (nm)
Red (660 nm)
600
700
800
900
1000
IR (940 nm)
Modulation Ratio (R)
SaO2 (%)
Red
20
40
60
80
100
IR
Red
IR
Red > IR
IR > Red
Shine Red and IR light through the skin and measure the absorption
DeoxyHb and OxyHb have different light absorption profiles
Using a standard curve, the ratio of Red/IR absorption (modulation ratio) is used to calculate the SaO2
It takes time for blood to flow from heart to skin, as a result pulse ox measurements lag by 5 – 15 seconds. Lag time may be shorter if probe is placed more centrally on forehead/ears, & longer if cardiac output is low. This is why SpO2 may continue to drop a few seconds after successful intubation.
ABNORMAL HEMOGLOBIN
· Methemoglobinemia (MetHb) causes a spurious reading, typically an SpO2 in the 85-88% range.
· Carboxyhemoglobinemia (CoHb) causes a false normal reading with SpO2 in the 94-100% range. Patients are hypoxic due to inability to unload O2 from Hb. (left shifted HbO2 curve)
· Sulfhemoglobin causes a spurious low SpO2 reading but patients may not be hypoxic (right shifted HbO2 curve)
· HbA1c > 7 may cause overestimation of SpO2, though the effect is usually small.
Cooximetry can measure MetHb & COHb
HR
SpO2
85
96
PI
0.8
Perfusion index indicates the signal strength
Amplitude may increase with greater SV (such as after a pause)
Oxygen Saturation value (SpO2); value is calculated from the curve and averaged over 3-6 sec
The plethysomograph waveform shows signal over time, corresponding to pulse
v1.0 (2020-12-17)
CC BY-SA 3.0
PROBE TYPE & POSITIONING
Pulse oximetry is the non-invasive continuous measurement of the oxygen saturation of hemoglobin. It exploits the fact that oxy- & deoxy hemoglobin differentially absorb red and infra-red light.
PRINCIPLE:
PLETHYSMOGRAPHY WAVEFORM ANALYSIS
MEDICATION EFFECTS
SKIN COLOR & NAIL POLISH
ACCURACY OF PULSE OXIMETRY:
Pulse oximetry may overestimate SpO2 in dark skinned individuals (compared with SaO2 on ABG), particularly when patients are hypoxemic. Black patients are 3x more likely to have occult hypoxemia.
Fingernail polish may decrease the accuracy of pulse oximetry, particularly darker colors (blue, green, black, & brown)
LOW FLOW STATES
CarboxyHb
The signal that varies with time represents the pulsatile blood flow.
Low perfusion states (such as due to high vasoconstriction or low cardiac output) can make the pulse oximetry signal weak or even undetectable. This can make monitoring SpO2 difficult or impossible in patients with non-pulsatile flow such as patients on ECMO or with an LVAD.
LOWER ACCURACY AT LOW SPO2
Because the calibration curve was developed using healthy volunteers, measured SpO2 can differ from SaO2 significantly at low values (e.g. an SpO2 less than 75%).
HYPEROXEMIA
Hyperoxia is harmful (particularly after cardiac arrest) but pulse oximetry cannot differentiate normal from supra-normal PaO2 if the SpO2 is 100%. Target SpO2 ≥94%
Methylene blue, isosulfan blue, fluorescein, & indocyanine green can transiently cause falsely low SpO2 readings.
CO-OXIMETRY
Instead of the two wavelengths used by standard pulse oximeters, co-oximetry uses 4 wavelengths to accurately measure abnormal hemoglobins.
PTT
Pulse transit time represents the interval from EKG R wave to PPG peak; determined by arterial compliance
vasoconstricted
· ↑ diastolic peak
· Shorter PTT, shorter ΔT
· ↓ perfusion index
normal
vasodilated
· ↑ systolic peak,↓ diastolic peak
· Longer PTT, longer/undefined ΔT
· ↑ perfusion index
systole
diastole
The plethysmography waveform has systolic & diastolic components; examination can provide physiologic clues about vascular tone/compliance.
Several factors can affect the accuracy of pulse oximetry readings, including:
Photodiode
ΔT
venous pulsations (severe TR)
· Late/prominent diastolic peak
· upsloping diastolic phase
Systolic peak
Diastolic peak
Notes: Nice review https://ccforum.biomedcentral.com/articles/10.1186/s13054-015-0984-8
Fingers best accuracy https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5905124/
Ears and forehead may be less effected by vasoconstriction. https://www.sciencedirect.com/science/article/pii/S095461111300053X
Effect of skin color https://pubs.asahq.org/anesthesiology/article/102/4/715/7364/Effects-of-Skin-Pigmentation-on-Pulse-Oximeter
Fingernail polish
https://journal.chestnet.org/article/S0012-3692(16)34864-4/fulltext
https://pubmed.ncbi.nlm.nih.gov/3382042/
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC6298168/
https://www.sciencedirect.com/science/article/abs/pii/S0964339713000864
No flow states
Pulseless oximetry (compress/release technique) https://pubmed.ncbi.nlm.nih.gov/26204107/