## Slide 1
percutaneous axial pumps (lv impella)
onepagericu.com
@nickmmark
@timbalthazar
Link to the most current version →
ONE
v1.0 (2023-05-29) CC BY-SA 3.0
Bleeding (33%) frequent complication at insertion site. More common if heparin used as purge solution (there is also a risk of HIT); consider using bicarbonate or dextrose.
Hemolysis (8%) may occur due to malpositioning or insufficient purge flow. Plasma free hemoglobin >40 mg/dL is most specific test. Reposition & ensure sufficient preload Severe cases can be treated with plasmapheresis or other therapies.
Limb ischemia – Obstruction of the femoral artery due to cannula without a re-prefusion canula (like ECMO); use NIR sensors on extremities to hasten diagnosis.
Suction events – inadequate LV filling or misplacement can cause suction to collapse the LV. This limits flow and can be arrhythmogenic; assess placement.
Cardiac Arrest – change to P2 (recommended) however can use higher program if confident about placement (see Cardiac Arrest on MCS OnePager)
COMPLICATIONS:
By continuously unloading the LV, an Impella reduces cardiac work. The degree of ventricular offloading depends on the level of support.
PROGRAM
PLACEMENT SIGNALS
Sheath sleeve allows sterile repositioning
A percutaneous axial flow pump that uses an impeller to move blood from the LV to the proximal aorta. This augments aortic and coronary pressures.
Forward flow depends on the model used (2.5, CP, LD, and 5.5) and the Program Setting. A purge fluid is added to avoid RBCs being drawn into the motor housing.
An Impella can provide more hemodynamic support than IABP. It can either be used for support during high-risk PCI or for hemodynamic support in ICU patients with cardiogenic shock. It can also be a vent in peripheral VA ECMO
PRINCIPLE:
PHYSIOLOGY:
Physiologically, an Impella functions as a temporary percutaneous LVAD (see LVAD OnePager), that continuously unloads the LV. This both reduces the native CO (reducing cardiac work & myocardial O2 consumption) and provides additional CO via the pump. This provides higher CO with lower cardiac work. Impella is preload dependent & afterload sensitive. Unlike IABP, Impella does not require EKG/pressure triggering so it is more tolerant of arrythmias.
Program level determines how fast the impeller rotates (RPMs), which is proportional to flow. Higher ”P level” corresponds to greater hemodynamic support.
There are two pressure waveforms used as placement signals. The Ao pressure is optically transduced at the outflow (Ao) positions, the inlet pressure (LV) is not measured but is calculated using the motor current and Ao pressures.
PLACEMENT:
Proper placement is essential. The device is typically placed under Fluoroscopic guidance ± TEE. POCUS/TTE are often used to verify placement. CXR can provide some info but does not show rotational malposition. Because the pump can become malpositioned, understanding & interpreting placement signals is necessary:
AORTIC
VALVE
OPENS
LV volume
LV pressure
UNASSISTED
IMPELLA (partial support; P3)
IMPELLA (full support; P9)
AORTIC VALVE CLOSES
LVSP = Systolic BP
LVEDP
Impeller pump is placed across the Aortic Valve; draws blood out of the LV
PURGE SOLUTION
MOTOR CURRENT
Motor current is proportional to pump torque and should increase with greater support (e.g. higher P-level). In patients with any native ventricular function, the motor current varies during the cardiac cycle.
A continuous infusion prevents blood from entering the pump motor housing. Dextrose solutions may contain Heparin or bicarbonate as additives. The goal is to achieve sufficient Purge Pressure to prevent RBCs entry. Can adjust the purge fluid and purge flow rate to achieve this.
Sheath side port can be used for blood
draws
Pressure sensor measures aortic pressure
Ao 94/60mmHg
584/496 mA
Impella Flow
Purge System
Cardiac Output
Cardiac Power
Output
P8
5.0 L/min
MENU
FLOW
PURGE
DISPLAY
MUTE
Placement Signal
Motor Current
6.0 L/min
1.0 watts
Purge Flow
Purge Pressure
LV 99/3 mmHg
PULSATILE
NON-PULSATILE
MOTOR CURRENT
D5W Dextrose 5% solution USP 1000mL
Lorem ipsum dolor sit amet, consectetur adipiscing elit, sed do eiusmod tempor incididunt ut labore et dolore magna aliqua. Ut enim ad minim veniam, quis nostrud exercitation ullamco laboris nisi ut aliquip ex ea commodo consequat. Duis aute irure dolor in reprehenderit in voluptate velit esse cillum dolore eu fugiat nulla pariatur. Excepteur sint occaecat cupidatat non proident, sunt in culpa qui officia deserunt mollit anim id est laborum.
Calculate Cardiac Output estimates the combined native and pump CO
AORTIC
DISPLACEMENT
DIASTOLIC
SUCTION
CONTINUOUS
SUCTION
CARDIAC
ARREST
• LVEDP <-40mmHg
• consider low preload
• LVEDP<-40mmhg
•Ao & LV tracings separated; consider malrotation/position
• normal Ao wave
• flat motor signal
VENTRICULAR
DISPLACEMENT
• wide swings in Ao pressure (LV pattern)
• flat motor signal
• lack of pulsatility in both motor signal & Ao pressure tracings
LV
Ao
NORMAL
SUCTION ALARM
LOW FLOW
ALARM
by Nick Mark MD &
Tim Balthazar MD & Jeff Scott DO
CONTRA-INDICATIONS:
Severe AS/AI (unable to pass catheter)
LV thrombus
Prosthetic Ao valve
Severe vascular disease
(unable to place catheter)